## calibration certificate
MAT |  | In-company calibration certificate | [6 empty] | RD-TU
Elab: JBELI.A_ 16.07.2025
Order number |  |  | 4000040566
Name |  |  | VT Hil 32
Equipement-ID |  |  | Marquardt
Manufacturer |  |  | RDT-TU
Type |  |  | ___
Serial no. |  |  | ___
Apadter no. |  |  | ___
Cost center |  |  | 203006
Location |  |  | MAT
Test sample |  |  | ___
Calibration intrucion |  |  | ___
This calibration certificate documents the traceability to national standards which realize the physical units of measurement according the International System of Units (SI). This calibration certificate may not be reproduced other than in full accept with the permission of the issuing laboratory.
Decision
Delivery condition |  |  | i.O. / ok | X | n.i.O. / not ok
Assessement | Serviceable | X | Conditional serviceable |  |  | - |  | not serviceable
Person in Charge | Dep. |  | RDT-TU |  |  |  | date of calibration |  |  | 09.06.2026
| Name: |  | Oussema Abidi |  |  |  | date of next calibration |  |  | 05-2027
Head of calibration dep. | Bennour S.
Note on E13: griffelm: Marquardt-Teilenummer auf Prüfling mit Profuktindex Marquardt-part number of the test sample with product index
Note on E14: griffelm: Marquardt-Teilenummer auf Prüfling mit Profuktindex Marquardt-part number of the test sample with product index
Note on A22: Is the sample capable for a calibration? Is the sample in a condition and can it be used as a golden or production sample; the meaning is here the acutal condition of the sample
Note on D24: "X" if appropriate
Note on H24: "X" if appropriate
Note on L24: "X" if appropriate
Note on E25: aiglstorferj: bei interner Kalibrierung: Abteilungsbezeichnung with internal calibration: name of department

## settings
PM-Values
| 13,5 V |  |  |  | 28 V |  |  |  | planned current Strom | target current at 13.5 V | target current at 28 V bei 28 V | measured with equipment
| theoretical value | measured value | corrected Value | tolerance | calculated value | measured value | corrected value | tolerance
1 | #REF! | 13454.6914 | 13436.6129 | 2% | #REF! | 28043.1309 | 27964.7091 | 2% | 0.001 | #REF! | #REF! | #REF!
inspection eq. values for the voltage sources
| theoretical value | measured value | measured with equipment
Voltage measurement tool
Instrument | Unites | Type | SAP N°
keithley | Voltage /current | keysight U1241B | 400.0036.165
|  | °C | K | Humidity %
Temperature |  | 24 | 298.15
Humidity |  |  |  | 55

## calibration procedure
Calibration Procedure
1 | calibration of the inspection equipment | To do that you've to use a calibrated measuring instrument. You've to measure the exact value of all resistor including their connecting cables.
2 | perform the voltage source calibration | Preparation: To verify that there are no line interruptions, set both voltage sources on the VT cards to 1 V Set the voltage sources about the test software to the target voltage values under the sheets of eash VT card. Measure the voltage between Kl.30 and Kl.31 with a calibrated voltage measuring instrument and put the measured value in the sheets . The validation will be done automatically.
3 | perform the voltage measurement calibration | Calibration is based on a target voltage set in the test software and a subsequent comparison with the measured voltage of the voltage channel to be calibrated. To do this, you need to connect the output of the voltage source used to the input of the voltage channel to be calibrated. In addition, a calibrated voltage measuring instrument must also be connected to the input voltage channel. The two voltages, one from the channel read into the test software and the other from the calibrated voltage meter, should be noted under the sheet.
5 | perform the current source calibration | To calibrate the current source, you need to place a load of resistance on it and set the corresponding current in the test software. Then measure the voltage drop across the resistor using a calibrated measuring instrument. The voltage should be noted on the sheet. The actual current will be calculated automatically and compared with the target value.
4 | perform the current measurement calibration | For the calibration you’ve to impress a defined current measure this current and validate this current against the measured current of the test system (test software). The impression of the current will be done about different resistors. These resistors had to be connected between the respective voltage source and the respective current measurement who had to be calibrated. Now you’ve to measure with a calibrated measuring instrument on the resistor the voltage drop. The voltage had to be written down under the sheet "current meas.
6 | perform the digital output checks | The checks of the digital outputs are only to be sure if they are working well. The outputs are digital outputs which will be set about the test software. You measure the levels them without a load on them and you confirm this with a “X” in the sheet “digital output”.
7 | Calibration of the equipment | The VT calibration is done by the System Calibration Manager is a separate program that is installed together with CANoe. It can be used to manage the adjustment

## System Calibration manager
Calibration Procedure
1 | calibration | In general, the calibration describes the process of determining the deviations of a physical quantity with regard to a reference. For the VT System modules, calibration means the comparison of the current values of the physical quantity under consideration with the accuracy ranges of the VT System manual. This will show with sufficient confidence whether a VT System module complies with the specified accuracy values or not
2 | Adjustment | In general, an adjustment is considered to be an intervention on the device where a mechanical or electrical modification or a modification via software is carried out to reduce deviations. For VT system modules, adjustment means storing correction factors so that the new current value of a physical quantity remains within the specified deviation range for the duration of the operation. The new current value of a physical quantity remains within the specified deviation range throughout the entire range of values in the VT system manual. value range of the VT system manual. If a VT channel is within the specification after calibration, an adjustment is not necessary. If a VT channel is outside the specification after calibration, an adjustment must be made.
3 | the System Calibration Manager | The VT System Calibration Manager is a stand-alone utility that is installed with CANoe. It allows to manage the setting of the VT system modules and to perform a self-test for specific VT system modules. This program can be launched from the CANoe menu detailed below below: Hardware ¦ Tools ¦ VT System Calibration Manager. In the same way as CANoe, the VT System Calibration Manager uses an ETH interface for an exclusive access to the VT System. The LAN cable used, must be connected directly to the backplane. After a scan of the VT System, the adjustable VT System modules in the tree view of the VT System Calibration Manager are shown in black. VT System modules without adjustment possibilities are shown in grey.
| External calibration | With the help of the external calibration, a test can be performed to show to what extent a VT System module and its values comply with the specified maximum deviations of the VT System manual. For this, a specific and limited wiring loom must be established as documented in the document VTSystem_Calibration_TestSpecification_EN.pdf. The document can be found in the CANoe installation directory, in the Doc subfolder. Below, an excerpt of this document is shown in order to explain the measurement setup for a channel. The external calibration is also supported by the document VTSystem_Calibration_TestProtocol.xls which provides a separate spreadsheet for every adjustable VT System module where measuring points are defined and where measured values can be entered. The table provides, through stored tests, the information on the compliance with the specified deviation or on the possibly required adjustment for the tested VT System module. If it is necessary to adjust a VT channel (VT System module), the export of the current correction parameters from the VT System module is required as a first step. Afterwards, the values can be saved as XML file with the extension *.vtparam by the VT System Calibration Manager.
4 | calculation of the factors | For the calculation of the new correction factors, new measurements have to be performed; the existing wiring loom can be used for this. As soon as new measuring values are entered, the correction parameters are recalculated.
6 | Calibrate the Device | After having measured the channels of a VT System module, the new correction parameters can be stored in an XML file The following import of the XML file by the VT System Calibration Manager allows the parameter transfer into the VT System module (see figure: Writing of new correction parameters). The effectiveness of the measures must be ensured by a new calibration. If the measure is successful, the values are within the specified deviation range.
7 | Self adjustment | Some of the VT System modules are provided with additional circuits to perform a self-adjustment. For this, the external wiring on the respective VT System module must be removed before starting the test. With the help of the VT System Calibration Manager, the self-adjustment which automatically writes new correction parameters into the corresponding VT System module can be started.

## VT 7001 1
measuring chain voltage source VT-7001 without load supp1 | [6 empty] | check chain voltage detected by source VT-7001 without load OUT1
target value | tolerance | valid area |  | Ubatt |  |  | target value | tolerance | valid area |  | Ubatt
|  | lower limit | upper limit | test system value | i.O. | n.i.O. |  |  | lower limit | upper limit | test system value | i.O. | n.i.O.
0 | 0.002% | -0.1 | 0.1 | -0.01099 | X |  | -2.375 | 0.001% | -2.495 | -2.275 | -2.3706 | X
5 | 0.002% | 4.8999 | 5.1001 | 4.9879 | X |  | -2 | 0.001% | -2.12 | -1.9 | -1.995 | X
10.003 | 0.002% | 9.9028 | 10.1032 | 9.9918 | X |  | -1.75 | 0.001% | -1.87 | -1.65 | -1.7449 | X
15 | 0.002% | 14.8997 | 15.1003 | 14.9785 | X |  | -1.5 | 0.001% | -1.62 | -1.4 | -1.4963 | X
20 | 0.002% | 19.8996 | 20.1004 | 19.9896 | X |  | -1.25 | 0.001% | -1.37 | -1.15 | -1.2459 | X
25 | 0.002% | 24.8995 | 25.1005 | 24.9776 | X |  | -1 | 0.001% | -1.12 | -0.90001 | -0.99592 | X
30 | 0.002% | 29.8994 | 30.1006 | 29.9789 | X |  | -0.75 | 0.001% | -0.86999 | -0.65001 | -0.7458 | X
35 | 0.002% | 34.8993 | 35.1007 | 34.9803 | X |  | -0.5 | 0.001% | -0.62 | -0.40001 | -0.49578 | X
measuring chain voltage source VT-7001 without load Supp2 | [6 empty] | check chain voltage detected by source VT-7001 without load OUT 2
target value | tolerance | valid area |  | Ubatt |  |  | target value | tolerance | valid area |  | Ubatt
|  | lower limit | upper limit | test system value | i.O. | n.i.O. |  |  | lower limit | upper limit | test system value | i.O. | n.i.O.
0 | 0.002% | -0.1 | 0.1 | -0.01221 | X |  | -2.375 | 0.001% | -2.495 | -2.275 | -2.3677 | X
5 | 0.002% | 4.8999 | 5.1001 | 4.9855 | X |  | -2 | 0.001% | -2.12 | -1.9 | -1.9918 | X
10 | 0.002% | 9.8998 | 10.1002 | 9.993 | X |  | -1.75 | 0.001% | -1.87 | -1.65 | -1.7419 | X
15 | 0.002% | 14.8997 | 15.1003 | 14.9956 | X |  | -1.5 | 0.001% | -1.62 | -1.4 | -1.4928 | X
20 | 0.002% | 19.8996 | 20.1004 | 19.9908 | X |  | -1.25 | 0.001% | -1.37 | -1.15 | -1.2424 | X
25 | 0.002% | 24.8995 | 25.1005 | 24.9959 | X |  | -1 | 0.001% | -1.12 | -0.90001 | -0.99248 | X
30 | 0.002% | 29.8994 | 30.1006 | 29.9863 | X |  | -0.75 | 0.001% | -0.86999 | -0.65001 | -0.74237 | X
35 | 0.002% | 34.8993 | 35.1007 | 34.9913 | X |  | -0.5 | 0.001% | -0.62 | -0.40001 | -0.49213 | X
measuring chain voltage source VT-7001 without load internal
target value | tolerance | valid area |  | Ubatt
|  | lower limit | upper limit | test system value | i.O. | n.i.O.
9.982 | 0.002% | 9.8818 | 10.0822 | 9.9881 | X
12.478 | 0.002% | 12.3778 | 12.5783 | 12.4918 | X
14.9729 | 0.002% | 14.8726 | 15.0732 | 14.9992 | X
17.4681 | 0.002% | 17.3678 | 17.5685 | 17.5079 | X
19.9634 | 0.002% | 19.863 | 20.0638 | 20.0006 | X
22.4591 | 0.002% | 22.3586 | 22.5595 | 22.5056 | X
24.9552 | 0.002% | 24.8548 | 25.0557 | 25.0032 | X
27.4531 | 0.002% | 27.3525 | 27.5536 | 27.496 | X
29.9499 | 0.002% | 29.8493 | 30.0505 | 29.9948 | X
Back to the calibration procedure

## VT 7001 2
measuring chain voltage source VT-7001 without load Internal | [6 empty] | check chain voltage detected by source VT-7001 without load OUT1
target value | tolerance | valid area |  | Ubatt |  |  | target value | tolerance | valid area |  | Ubatt
|  | lower limit | upper limit | test system value | i.O. | n.i.O. |  |  | lower limit | upper limit | test system value | i.O. | n.i.O.
9.982 | 0.002% | 9.8818 | 10.0822 | 10.0028 | X |  | -2.375 | 0.001% | -2.475 | -2.275 | -2.3736 | X
12.478 | 0.002% | 12.3778 | 12.5783 | 12.5028 | X |  | -2 | 0.001% | -2.1 | -1.9 | -1.9976 | X
14.9729 | 0.002% | 14.8726 | 15.0732 | 15.0029 | X |  | -1.75 | 0.001% | -1.85 | -1.65 | -1.7474 | X
17.4681 | 0.002% | 17.3678 | 17.5685 | 17.5005 | X |  | -1.5 | 0.001% | -1.6 | -1.4 | -1.4983 | X
19.9634 | 0.002% | 19.863 | 20.0638 | 19.9957 | X |  | -1.25 | 0.001% | -1.35 | -1.15 | -1.2478 | X
22.4591 | 0.002% | 22.3586 | 22.5595 | 22.4982 | X |  | -1 | 0.001% | -1.1 | -0.90001 | -0.99769 | X
24.9552 | 0.002% | 24.8548 | 25.0557 | 24.9983 | X
27.4531 | 0.002% | 27.3525 | 27.5536 | 27.5033 | X
29.9499 | 0.002% | 29.8493 | 30.0505 | 29.9985 | X
measuring chain voltage source VT-7001 without load Supp1 | [6 empty] | check chain voltage detected by source VT-7001 without load OUT 2
target value | tolerance | valid area |  | Ubatt |  |  | target value | tolerance | valid area |  | Ubatt
|  | lower limit | upper limit | test system value | i.O. | n.i.O. |  |  | lower limit | upper limit | test system value | i.O. | n.i.O.
0 | 0.002% | -0.1 | 0.1 | -0.01221 | X |  | -2.375 | 0.001% | -2.475 | -2.275 | -2.3723 | X
5 | 0.002% | 4.8999 | 5.1001 | 4.9916 | X |  | -2 | 0.001% | -2.1 | -1.9 | -1.9967 | X
10.003 | 0.002% | 9.9028 | 10.1032 | 10.0247 | X |  | -1.75 | 0.001% | -1.85 | -1.65 | -1.7467 | X
15 | 0.002% | 14.8997 | 15.1003 | 15.0395 | X |  | -1.5 | 0.001% | -1.6 | -1.4 | -1.4978 | X
20 | 0.002% | 19.8996 | 20.1004 | 20.0421 | X |  | -1.25 | 0.001% | -1.35 | -1.15 | -1.2474 | X
25 | 0.002% | 24.8995 | 25.1005 | 25.063 | X |  | -1 | 0.001% | -1.1 | -0.90001 | -0.99751 | X
30 | 0.002% | 29.8994 | 30.1006 | 30.0827 | X
measuring chain voltage source VT-7001 without load Supp2
target value | tolerance | valid area |  | Ubatt
|  | lower limit | upper limit | test system value | i.O. | n.i.O.
0 | 0.002% | -0.1 | 0.1 | -0.03662 | X
15 | 0.002% | 14.8997 | 15.1003 | 15.0493 | X
20 | 0.002% | 19.8996 | 20.1004 | 20.0665 | X
25 | 0.002% | 24.8995 | 25.1005 | 25.0826 | X
30 | 0.002% | 29.8994 | 30.1006 | 30.0998 | X
35 | 0.002% | 34.8993 | 35.1007 | 35.0998 | X
Back to the calibration procedure

## VT 1004
measuring chain voltage VT 1004 1
planned voltage | tolerance | Input 1 |  |  |  | Input 2 |  |  |  | Input 3 |  |  |  | Input 4
|  | test system value | reference Value | i.O. | n.i.O. | test system value | reference Value | i.O. | n.i.O. | test system value | reference Value | i.O. | n.i.O. | test system value | reference Value | i.O. | n.i.O.
[V] | [%] | [V] | [V]
0 | 1.2% | 0 | 0 | X |  | 0 | 0 | X |  | 0 | 0 | X |  | 0 | 0 | X
5 | 1.2% | 4.9916 | 5 | X |  | 4.994 | 5 | X |  | 5.0014 | 5 | X |  | 4.994 | 5 | X
10 | 1.2% | 9.9966 | 10 | X |  | 9.9966 | 10 | X |  | 9.9991 | 10 | X |  | 9.9954 | 10 | X
15 | 1.2% | 15.0041 | 15 | X |  | 14.998 | 15 | X |  | 14.9846 | 15 | X |  | 14.9858 | 15 | X
20 | 1.2% | 20.0018 | 20 | X |  | 19.9957 | 20 | X |  | 19.9969 | 20 | X |  | 19.9994 | 20 | X
25 | 1.2% | 25.0093 | 25 | X |  | 24.9971 | 25 | X |  | 25.0044 | 25 | X |  | 24.9983 | 25 | X
30 | 1.2% | 29.9985 | 30 | X |  | 29.9753 | 30 | X |  | 30.018 | 30 | X |  | 29.9924 | 30 | X
35 | 1.2% | 35.0047 | 35 | X |  | 35.0108 | 35 | X |  | 34.9925 | 35 | X |  | 35.0108 | 35 | X
measuring chain voltage VT 1004 2
planned voltage | tolerance | Input 1 |  |  |  | Input 2 |  |  |  | Input 3 |  |  |  | Input 4
|  | test system value | reference Value | i.O. | n.i.O. | test system value | reference Value | i.O. | n.i.O. | test system value | reference Value | i.O. | n.i.O. | test system value | reference Value | i.O. | n.i.O.
0 | 1.2% | 0 | 0 | X |  | 0 | 0 | X |  | 0 | 0 | X |  | 0 | 0 | X
5 | 1.2% | 4.9904 | 5 | X |  | 4.9892 | 5 | X |  | 5.0026 | 5 | X |  | 5.0001 | 5 | X
10 | 1.2% | 10.0064 | 10 | X |  | 9.993 | 10 | X |  | 9.9979 | 10 | X |  | 10.0028 | 10 | X
15 | 1.2% | 14.9541 | 15 | X |  | 14.9944 | 15 | X |  | 14.9907 | 15 | X |  | 14.9931 | 15 | X
20 | 1.2% | 20.0079 | 20 | X |  | 19.9994 | 20 | X |  | 19.9957 | 20 | X |  | 19.9969 | 20 | X
25 | 1.2% | 24.9995 | 25 | X |  | 25.0032 | 25 | X |  | 25.0093 | 25 | X |  | 25.0056 | 25 | X
30 | 1.2% | 30.0046 | 30 | X |  | 30.0058 | 30 | X |  | 30.0058 | 30 | X |  | 30.119 | 30 | X
35 | 1.2% | 35.0059 | 35 | X |  | 35.0011 | 35 | X |  | 35.0267 | 35 | X |  | 35.0145 | 35 | X
measuring chain voltage VT 1004 3
planned voltage | tolerance | Input 1 |  |  |  | Input 2 |  |  |  | Input 3 |  |  |  | Input 4
|  | test system value | reference Value | i.O. | n.i.O. | test system value | reference Value | i.O. | n.i.O. | test system value | reference Value | i.O. | n.i.O. | test system value | reference Value | i.O. | n.i.O.
0 | 1.2% | 0 | 0 | X |  | 0 | 0 | X |  | 0 | 0 | X |  | 0 | 0 | X
5 | 1.2% | 4.9977 | 5 | X |  | 4.9953 | 5 | X |  | 5.0014 | 5 | X |  | 4.9818 | 5 | X
10 | 1.2% | 9.9979 | 10 | X |  | 10.0064 | 10 | X |  | 9.9954 | 10 | X |  | 9.9869 | 10 | X
15 | 1.2% | 15.0066 | 15 | X |  | 14.9919 | 15 | X |  | 14.9883 | 15 | X |  | 15.0029 | 15 | X
20 | 1.2% | 20.0031 | 20 | X |  | 20.0104 | 20 | X |  | 20.0018 | 20 | X |  | 19.9799 | 20 | X
25 | 1.2% | 24.9971 | 25 | X |  | 25.002 | 25 | X |  | 25.0032 | 25 | X |  | 24.9886 | 25 | X
30 | 1.2% | 29.996 | 30 | X |  | 30.0278 | 30 | X |  | 29.9948 | 30 | X |  | 30.0143 | 30 | X
35 | 1.2% | 35.0157 | 35 | X |  | 34.9986 | 35 | X |  | 34.9937 | 35 | X |  | 34.9974 | 35 | X

## VT 2004
measuring chain voltage VT 2004 1
planned voltage | tolerance | Input 1 |  |  |  | Input 2 |  |  |  | Input 3 |  |  |  | Input 4
|  | test system value | reference Value | i.O. | n.i.O. | test system value | reference Value | i.O. | n.i.O. | test system value | reference Value | i.O. | n.i.O. | test system value | reference Value | i.O. | n.i.O.
[V] | [%] | [V] | [V]
1 | 0.05% | 1.0004 | 1 | X |  | 3.001 | 3 | X |  | 3.001 | 3 | X |  | 3.999 | 4 | X
5 | 0.05% | 4.9979 | 5 | X |  | 4.999 | 5 | X |  | 4.999 | 5 | X |  | 4.999 | 5 | X
10 | 0.05% | 9.9982 | 10 | X |  | 5.99 | 5.99 | X |  | 6.001 | 5.999 | X |  | 5.99 | 5.99 | X
15 | 0.05% | 14.9959 | 15 | X |  | 10.001 | 10 | X |  | 10.001 | 10 | X |  | 10.001 | 10 | X
20 | 0.05% | 19.9962 | 20 | X
measuring chain voltage VT 2004 2
planned voltage | tolerance | Input 1 |  |  |  | Input 2 |  |  |  | Input 3 |  |  |  | Input 4
|  | test system value | reference Value | i.O. | n.i.O. | test system value | reference Value | i.O. | n.i.O. | test system value | reference Value | i.O. | n.i.O. | test system value | reference Value | i.O. | n.i.O.
10 | 0.05% | 9.999 | 10 | X |  | 9.998 | 10 | X |  | 9.999 | 10 | X |  | 10 | 10 | X
6 | 0.05% | 6.001 | 6 | X |  | 6 | 6 | X |  | 6 | 6 | X |  | 6 | 6 | X
5 | 0.05% | 5.001 | 5 | X |  | 5 | 5 | X |  | 5 | 5 | X |  | 5.001 | 5 | X
4 | 0.05% | 4 | 4 | X |  | 4 | 4 | X |  | 4 | 4 | X |  | 4.001 | 4 | X
0 | 0.05% | 0 | 0 | X |  | 0 | 0 | X |  | 0 | 0 | X |  | 0 | 0 | X
measuring chain voltage VT 2004 3
planned voltage | tolerance | Input 1 |  |  |  | Input 2 |  |  |  | Input 3 |  |  |  | Input 4
|  | test system value | reference Value | i.O. | n.i.O. | test system value | reference Value | i.O. | n.i.O. | test system value | reference Value | i.O. | n.i.O. | test system value | reference Value | i.O. | n.i.O.
9 | 0.05% | 9 | 9 | X |  | 9 | 9 | X |  | 9 | 9 | X |  | 9 | 9 | X
7 | 0.05% | 6.999 | 7 | X |  | 7.001 | 7 | X |  | 6.997 | 7 | X |  | 6.998 | 7 | X
6 | 0.05% | 5.999 | 6 | X |  | 6.001 | 6 | X |  | 6 | 6 | X |  | 5.997 | 6 | X
5 | 0.05% | 5.001 | 5 | X |  | 5 | 5 | X |  | 5.002 | 5 | X |  | 5 | 5 | X
measuring chain voltage VT 2004 4
planned voltage | tolerance | Input 1 |  |  |  | Input 2 |  |  |  | Input 3 |  |  |  | Input 4
|  | test system value | reference Value | i.O. | n.i.O. | test system value | reference Value | i.O. | n.i.O. | test system value | reference Value | i.O. | n.i.O. | test system value | reference Value | i.O. | n.i.O.
9 | 0.05% | 9 | 9 | X |  | 8.999 | 9 | X |  | 9 | 9 | X |  | 9 | 9 | X
7 | 0.05% | 7 | 7 | X |  | 6.999 | 7 | X |  | 6.998 | 7 | X |  | 7 | 7 | X
6 | 0.05% | 6 | 6 | X |  | 6.001 | 6 | X |  | 5.997 | 6 | X |  | 5.999 | 6 | X
5 | 0.05% | 5 | 5 | X |  | 5 | 5 | X |  | 5.002 | 5 | X |  | 5 | 5 | X

## VT 2516
measuring chain voltage VT 2516 1 | [36 empty] | measuring chain voltage VT 2516 1
planned voltage | tolerance | Input 1 |  |  |  | Input 2 |  |  |  | Input 3 |  |  |  | Input 4 |  |  |  | Input 5 |  |  |  | Input 6 |  |  |  | Input 7 |  |  |  | Input 8 |  |  |  | Input 9 |  |  |  | Input 10 |  |  |  | Input 11 |  |  |  | Input 12 |  |  |  | Input 13 |  |  |  | Input 14 |  |  |  | Input 15 |  |  |  | Input 16
|  | test system value | reference Value | i.O. | n.i.O. | test system value | reference Value | i.O. | n.i.O. | test system value | reference Value | i.O. | n.i.O. | test system value | reference Value | i.O. | n.i.O. | test system value | reference Value | i.O. | n.i.O. | test system value | reference Value | i.O. | n.i.O. | test system value | reference Value | i.O. | n.i.O. | test system value | reference Value | i.O. | n.i.O. | test system value | reference Value | i.O. | n.i.O. | test system value | reference Value | i.O. | n.i.O. | test system value | reference Value | i.O. | n.i.O. | test system value | reference Value | i.O. | n.i.O. | test system value | reference Value | i.O. | n.i.O. | test system value | reference Value | i.O. | n.i.O. | test system value | reference Value | i.O. | n.i.O. | test system value | reference Value | i.O. | n.i.O.
[V] | [%] | [V] | [V]
0 | 0.5% | 0 | 0 | X |  | 0 | 0 | X |  | 0 | 0 | X |  | 0 | 0 | X |  | 0 | 0 | X |  | 0 | 0 | X |  | 0 | 0 | X |  | 0 | 0 | X |  | 0 | 0 | X |  | 0 | 0 | X |  | 0 | 0 | X |  | 0 | 0 | X |  | 0 | 0 | X |  | 0 | 0 | X |  | 0 | 0 | X |  | 0 | 0 | X
5 | 0.5% | 5 | 5 | X |  | 4.98 | 5 | X |  | 4.98 | 5 | X |  | 4.98 | 5 | X |  | 5 | 5 | X |  | 4.98 | 5 | X |  | 4.98 | 5 | X |  | 4.98 | 5 | X |  | 4.98 | 5 | X |  | 4.98 | 5 | X |  | 4.98 | 5 | X |  | 4.98 | 5 | X |  | 4.98 | 5 | X |  | 4.98 | 5 | X |  | 4.98 | 5 | X |  | 4.98 | 5 | X
10 | 0.5% | 9.98 | 10 | X |  | 9.98 | 10 | X |  | 9.98 | 10 | X |  | 9.98 | 10 | X |  | 9.98 | 10 | X |  | 9.98 | 10 | X |  | 9.98 | 10 | X |  | 9.98 | 10 | X |  | 9.98 | 10 | X |  | 9.98 | 10 | X |  | 9.98 | 10 | X |  | 9.98 | 10 | X |  | 9.98 | 10 | X |  | 9.98 | 10 | X |  | 9.98 | 10 | X |  | 9.98 | 10 | X
15 | 0.5% | 14.98 | 15 | X |  | 14.98 | 15 | X |  | 14.96 | 15 | X |  | 14.98 | 15 | X |  | 14.98 | 15 | X |  | 14.96 | 15 | X |  | 14.96 | 15 | X |  | 14.96 | 15 | X |  | 14.96 | 15 | X |  | 14.96 | 15 | X |  | 14.96 | 15 | X |  | 14.96 | 15 | X |  | 14.96 | 15 | X |  | 14.96 | 15 | X |  | 14.96 | 15 | X |  | 14.96 | 15 | X
20 | 0.5% | 19.98 | 20 | X |  | 19.96 | 20 | X |  | 19.96 | 20 | X |  | 19.98 | 20 | X |  | 19.96 | 20 | X |  | 19.96 | 20 | X |  | 19.96 | 20 | X |  | 19.96 | 20 | X |  | 19.96 | 20 | X |  | 19.96 | 20 | X |  | 19.96 | 20 | X |  | 19.96 | 20 | X |  | 19.96 | 20 | X |  | 19.96 | 20 | X |  | 19.96 | 20 | X |  | 19.96 | 20 | X
25 | 0.5% | 24.96 | 25 | X |  | 24.96 | 25 | X |  | 24.96 | 25 | X |  | 24.96 | 25 | X |  | 24.96 | 25 | X |  | 24.96 | 25 | X |  | 24.96 | 25 | X |  | 24.96 | 25 | X |  | 24.96 | 25 | X |  | 24.96 | 25 | X |  | 24.96 | 25 | X |  | 24.96 | 25 | X |  | 24.96 | 25 | X |  | 24.96 | 25 | X |  | 24.96 | 25 | X |  | 24.96 | 25 | X
30 | 0.5% | 29.96 | 30 | X |  | 29.94 | 30 | X |  | 29.94 | 30 | X |  | 29.96 | 30 | X |  | 29.94 | 30 | X |  | 29.96 | 30 | X |  | 29.96 | 30 | X |  | 29.96 | 30 | X |  | 29.96 | 30 | X |  | 29.96 | 30 | X |  | 29.96 | 30 | X |  | 29.96 | 30 | X |  | 29.96 | 30 | X |  | 29.96 | 30 | X |  | 29.96 | 30 | X |  | 29.96 | 30 | X
35 | 0.5% | 34.96 | 35 | X |  | 34.94 | 35 | X |  | 34.94 | 35 | X |  | 34.94 | 35 | X |  | 34.94 | 35 | X |  | 34.94 | 35 | X |  | 34.94 | 35 | X |  | 34.94 | 35 | X |  | 34.94 | 35 | X |  | 34.94 | 35 | X |  | 34.94 | 35 | X |  | 34.94 | 35 | X |  | 34.94 | 35 | X |  | 34.94 | 35 | X |  | 34.94 | 35 | X |  | 34.94 | 35 | X
measuring chain voltage VT 2516 2 | [36 empty] | measuring chain voltage VT 2516 2
planned voltage | tolerance | Input 1 |  |  |  | Input 2 |  |  |  | Input 3 |  |  |  | Input 4 |  |  |  | Input 5 |  |  |  | Input 6 |  |  |  | Input 7 |  |  |  | Input 8 |  |  |  | Input 9 |  |  |  | Input 10 |  |  |  | Input 11 |  |  |  | Input 12 |  |  |  | Input 13 |  |  |  | Input 14 |  |  |  | Input 15 |  |  |  | Input 16
|  | test system value | reference Value | i.O. | n.i.O. | test system value | reference Value | i.O. | n.i.O. | test system value | reference Value | i.O. | n.i.O. | test system value | reference Value | i.O. | n.i.O. | test system value | reference Value | i.O. | n.i.O. | test system value | reference Value | i.O. | n.i.O. | test system value | reference Value | i.O. | n.i.O. | test system value | reference Value | i.O. | n.i.O. | test system value | reference Value | i.O. | n.i.O. | test system value | reference Value | i.O. | n.i.O. | test system value | reference Value | i.O. | n.i.O. | test system value | reference Value | i.O. | n.i.O. | test system value | reference Value | i.O. | n.i.O. | test system value | reference Value | i.O. | n.i.O. | test system value | reference Value | i.O. | n.i.O. | test system value | reference Value | i.O. | n.i.O.
[V] | [%] | [V] | [V]
9 | 0.5% | 9.022 | 9.01 | X |  | 9.022 | 9.01 | X |  | 3.514 | 3.509 | X |  | 1.678 | 1.67 | X |  | 6.858 | 6.86 | X |  | 6.386 | 6.398 | X |  | 6.366 | 6.369 | X |  | 6.346 | 6.36 | X |  | 6.346 | 6.356 | X |  | 6.444 | 6.438 | X |  | 6.424 | 6.45 | X |  | 6.444 | 6.447 | X |  | 6.346 | 6.36 | X |  | 6.366 | 6.371 | X |  | 6.346 | 6.366 | X |  | 6.386 | 6.397 | X
7 | 0.5% | 6.979 | 7 | X |  | 6.979 | 6.949 | X |  | 2.256 | 2.248 | X |  | 3.046 | 3.04 | X |  | 8.236 | 8.24 | X |  | 8.456 | 8.46 | X |  | 8.514 | 8.51 | X |  | 8.514 | 8.52 | X |  | 8.496 | 8.51 | X |  | 8.476 | 8.48 | X |  | 8.474 | 8.49 | X |  | 8.476 | 8.49 | X |  | 8.496 | 8.52 | X |  | 8.534 | 8.53 | X |  | 8.508 | 8.52 | X |  | 8.436 | 8.46 | X
6 | 0.5% | 5.996 | 5.98 | X |  | 5.996 | 5.98 | X |  | 6.992 | 7 | X |  | 6.97 | 6.97 | X |  | 9.492 | 9.5 | X |  | 12.558 | 12.58 | X |  | 12.636 | 12.63 | X |  | 12.578 | 12.6 | X |  | 12.578 | 12.6 | X |  | 12.558 | 12.57 | X |  | 12.578 | 12.61 | X |  | 12.594 | 12.6 | X |  | 12.596 | 12.62 | X |  | 12.558 | 12.56 | X |  | 12.596 | 12.61 | X |  | 12.558 | 12.58 | X
5 | 0.5% | 5 | 5 | X |  | 5.018 | 4.998 | X |  | 11.97 | 11.97 | X |  | 11.952 | 11.96 | X |  | 10.778 | 10.78 | X |  | 15.996 | 16 | X |  | 15.956 | 15.96 | X |  | 16.014 | 16.02 | X |  | 16.012 | 16.03 | X |  | 15.996 | 16 | X |  | 15.916 | 15.95 | X |  | 16.014 | 16.03 | X |  | 15.936 | 15.96 | X |  | 15.996 | 16 | X |  | 16.014 | 16.03 | X |  | 15.996 | 16.02 | X

## VT 2816
measuring chain voltage VT 2816 1
planned voltage | tolerance | Input 1 |  |  |  | Input 2 |  |  |  | Input 3 |  |  |  | Input 4 |  |  |  | Input 5 |  |  |  | Input 6 |  |  |  | Input 7 |  |  |  | Input 8 |  |  |  | Input 9 |  |  |  | Input 10 |  |  |  | Input 11 |  |  |  | Input 12 |  |  |  | Input 13 |  |  |  | Input 14 |  |  |  | Input 15 |  |  |  | Input 16 |  |  |  | Input 16
|  | test system value | reference Value | i.O. | n.i.O. | test system value | reference Value | i.O. | n.i.O. | test system value | reference Value | i.O. | n.i.O. | test system value | reference Value | i.O. | n.i.O. | test system value | reference Value | i.O. | n.i.O. | test system value | reference Value | i.O. | n.i.O. | test system value | reference Value | i.O. | n.i.O. | test system value | reference Value | i.O. | n.i.O. | test system value | reference Value | i.O. | n.i.O. | test system value | reference Value | i.O. | n.i.O. | test system value | reference Value | i.O. | n.i.O. | test system value | reference Value | i.O. | n.i.O. | test system value | reference Value | i.O. | n.i.O. | test system value | reference Value | i.O. | n.i.O. | test system value | reference Value | i.O. | n.i.O. | test system value | reference Value | i.O. | n.i.O. | test system value | reference Value | i.O. | n.i.O.
[V] | [%] | [V] | [V]
0.994 | 0.5% | 0.996 | 0.994 | X |  | 0.996 | 0.994 | X |  | 0.996 | 0.994 | X |  | 0.996 | 0.994 | X |  | 0.996 | 0.994 | X |  | 0.996 | 0.994 | X |  | 0.996 | 0.994 | X |  | 0.996 | 0.994 | X |  | 0.996 | 0.994 | X |  | 0.996 | 0.994 | X |  | 0.996 | 0.994 | X |  | 0.996 | 0.994 | X |  | 0.996 | 0.994 | X |  | 12 | 12 | X |  | 12 | 12 | X |  | 12 | 12 | X |  | 12 | 12 | X
1.994 | 0.5% | 1.996 | 1.994 | X |  | 1.996 | 1.994 | X |  | 1.996 | 1.994 | X |  | 1.996 | 1.994 | X |  | 1.996 | 1.994 | X |  | 1.996 | 1.994 | X |  | 1.996 | 1.994 | X |  | 1.996 | 1.994 | X |  | 1.996 | 1.994 | X |  | 1.996 | 1.994 | X |  | 1.996 | 1.994 | X |  | 1.996 | 1.994 | X |  | 1.996 | 1.994 | X |  | 10 | 10 | X |  | 10 | 10 | X |  | 10 | 10 | X |  | 10 | 10 | X
2.996 | 0.5% | 2.996 | 2.996 | X |  | 2.996 | 2.996 | X |  | 2.996 | 2.996 | X |  | 2.996 | 2.996 | X |  | 2.996 | 2.996 | X |  | 2.996 | 2.996 | X |  | 2.996 | 2.996 | X |  | 2.996 | 2.996 | X |  | 2.996 | 2.996 | X |  | 2.996 | 2.996 | X |  | 2.996 | 2.996 | X |  | 2.996 | 2.996 | X |  | 2.996 | 2.996 | X |  | 6 | 6 | X |  | 6 | 6 | X |  | 6 | 6 | X |  | 6 | 6 | X
9.997 | 0.5% | 10 | 9.997 | X |  | 10 | 9.997 | X |  | 10 | 9.997 | X |  | 10 | 9.997 | X |  | 10 | 9.997 | X |  | 10 | 9.997 | X |  | 10 | 9.997 | X |  | 10 | 9.997 | X |  | 10 | 9.997 | X |  | 10 | 9.997 | X |  | 10 | 9.997 | X |  | 10 | 9.997 | X |  | 10 | 9.997 | X |  | 4.998 | 5 | X |  | 5 | 5 | X |  | 4.998 | 5 | X |  | 5 | 5 | X